## Slide 1: Vbrain sequences
Dec 7, 2022

## Slide 2: What is Vbrain?
Initiative by Dr. Sophia Frangou to create an image repository for scanning at UBC
Will allow the combination of studies to create larger datasets

## Slide 3: Recommended sequences
In order to be able to combine datasets, there needs to be similar sequences collected
The Centre has put together recommended sequences
  If data collected with a Vbrain sequence, then it can go in the registry

## Slide 4: 3DT1 - MPRAGE
 | Philips MPRAGE | Philips MPRAGE | GE MPRAGE | GE MPRAGE
 | Philips High Res | Philips | GE PETMR | GE MR Only
Coil | dSTREAM 32 channel | dSTREAM 32 channel | HNU 12 channel | HNU 12 channel
Orientation | Sagittal Oblique | Sagittal Oblique | Axial | Axial/Axial Oblique
FOV (mm) | 256 | 256 | 256 | 256
Slice thickness (mm) | 0.8 | 1 | 1 | 1
Slice gap (mm) | NA | NA | NA | NA
#Slices | 225 | 180 | 206 | 206
Slice coverage (mm) | 180 | 180 | 206 | 206
Matrix | 320 | 256 | 256 | 256
Voxel size FxPxS (mm) | 0.8x0.8x0.8 | 1x1x1 | 1x1x1 | 1x1x1
TR (ms) | 9.3 | 7.7 | 8.4 | 8.4
TE (ms) | 4.3 | 3.6 | 3.2 | 3.2
Shot Interval (ms) | 2400 | 2400 | 2783 | 2783
Inversion Time | 950 | 950 | 900 | 900
Flip angle (deg) | 8 | 8 | 8 | 8
Parallel Imaging Factor | CS = 3.6 | CS = 3.6 | Phase=2, Slice = 1 | Phase=2, Slice = 1
Acquisition time | 5:50 | 3:42 | 5:39 | 5:39

## Slide 5: Resting state fMRI
 | Philips rsfMRI | GE rsfMRI | GE rsfMRI
 | Philips | GE PETMR | GE MR only
Coil | dSTREAM 32 channel | HNU 12 channel | HNU 12 channel
Orientation | Axial Oblique | Axial(PETMR) | Axial/Axial oblique(MR only)
FOV (mm) | 240 | 258 | 258
Slice thickness (mm) | 2.5 | 3 | 3
Slice gap (mm) | 0 | 0 | 0
#Slices | 62 | 62 | 62
Slice coverage (mm) | 155 | 186 | 186
Matrix | 96 | 86 | 86
Voxel size FxPxS (mm) | 2.5x2.5x2.5 | 3x3x3 | 3x3x3
TR (ms) | 2000 | 2000 | 2000
TE (ms) | 30 | 25 | 25
Flip angle (deg) | 77 | 77 | 77
Echoes | 1 | 1 | 1
Parallel Imaging Factor | 2 | 2 | 2
Multiband factor | 2 | 2 | 2
Acquisition time | 8-15 min | 8-15 min | 8-15 min

## Slide 6: Diffusion Tensor Imaging
 | Philips DTI | Philips DTI | Philips DTI | Philips DTI | GE DTI | GE DTI | GE DTI
 | Philips Hardi60 b700 | Philips Hardi60 b1000 | Philips 98dir 3b | Philips 128dir 3b | GE Hardi60 | GE NODDI | GE DBSI
Coil | 32 channel | 32 channel | 32 channel | 32 channel | HNU 12 channel | HNU 12 channel | HNU 12 channel
Orientation | Axial Oblique | Axial Oblique | Axial Oblique | Axial Oblique | Axial | Axial | Axial
FOV (mm) | 224 | 224 | 224 | 224 | 224 | 224 | 224
Slice thickness (mm) | 2.2 | 2.2 | 2 | 2 | 2.2 | 2.2 | 2.2
Slice gap (mm) | 0 | 0 | 0 | 0 | NA | NA | NA
#Slices | 70 | 70 | 75 | 75 | 80 | 80 | 80
Slice coverage (mm) | 154 | 154 | 150 | 150 | 176 | 176 | 176
Acquisition Matrix | 104x102 | 104x102 | 112 | 112 | 100x100 | 100x100 | 100x100
Voxel size FxPxS (mm) | 2.2x2.2x2.2 | 2.2x2.2x2.2 | 2x2x2 | 2x2x2 | 2.2x2.2x2.2 | 2.2x2.2x2.2 | 2.2x2.2x2.2
Reconstruction matrix | 112 | 112 | 112 | 112 | 256x256 | 256x256 | 256x256
TR (ms) | 5266 | 5266 | 4450 | 4250 | 8193 | 7576 | 7465
TE (ms) | 68 | 68 | 95 | 95 | 69.7 | 88.7 | 82.7
b values(s/mm2) | 0, 700 | 0, 1000 | 0, 500, 1000, 2000 | 0, 500, 1000, 2000 | 0, 700 | 9 values (0-2000) | 8 values (0-1500)
number of directions | 60 | 60 | 98 | 128 | 60 | 106 | 98
Number of T2 images | 3 | 3 | 10 | 12 | 5 | 3 | 3
Multiband factor | 2 | 2 | 3 | 3 | 2 | 2 | 2
Parallel Imaging Factor | 2.1 | 2.1 | 1 | 1 | Phase=2, Slice=1 | Phase=2, Slice=1 | Phase=2, Slice=1
Acquisition time | 5:32 | 5:32 | 7:45 | 10:01 | 9:01 | 13:53 | 12:41
Reverse phase | 0:11 | 0:11 | 0:13 | 0:13 | 1:20 | 1:26 | 1:21
Total scan time | 5:43 | 5:43 | 7:58 | 10:14 | 10:21 | 15:19 | 14:02

## Slide 7: Susceptibility Weighted Imaging
 | Philips High Res SWI | Philips fast SWI | GE SWI
 | Philips | Philips | GE SWI
Coil | 32 ch | 32 ch | HNU 12 channel
Orientation | Axial | Axial | Axial
FOV (mm) | 220 x 180 x 170 | 220 x 180 x 170 | 220x176
Slice thickness (mm) | 0.9 acq, 076 rec | 1.3 mm acq; 0.76 rec | 2
Slice gap (mm) | N/A | N/A | NA
#Slices | 224 | 224 | 90
Slice coverage (mm) | 170 mm | 180mm | 180
Matrix |  | 288 | 440x288
Voxel size FxPxS (mm) | 0.9 x 0.9 x 0.9, rec: 0.76 x 0.76 x 0.76 | 0.8 x 1.1 x 1.3, rec: 0.76 x 0.76 x 0.76 | 0.5x0.6x2
BW (kHz) | 180 Hz | 180Hz | 83.33
TE (ms), TR | 5.2, Delta TE = 6.4, TR = 35 | 5.2, Delta TE = 6.4, TR = 35 | 5, 10.1, 15.1, 20.2, 25.2
number of echoes | 5 | 5 | 5
Flip angle (deg) | 15 | 15 | 20
Parallel Imaging Factor | CS = 4.5 | CS = 6 | Phase=2, Slice = 1
Acquisition time | 5:37 | 2:26 | 8:12
scan name | VBRAIN_QSM_0_9_FLOW_COMP | VBRAIN_QSM_08_1_2_CS6_FLOW_COMP | 
flyback | yes | yes | 
flow comp | yes | yes | 

## Slide 8: Magnetic Resonance Spectroscopy
 | Philips Elition semiLASER | GE PET/MR semiLASER
 | Philips | GE PET/MR
Coil | dSTREAM 32 channel | HNU 12 channel
Orientation | Study Specific | Study Specific
Minimum Voxel Volume | 6 mL | 0.14mL (Recommend>=6mL)
Minimum Shim Volume | 2cm in each direction | Shim Volume is the sphere around the voxel
Samples | 2048 | 4096
Spectral Bandwidth | 2000 | 5000
Centre Frequency | 2.0 ppm | 2.7ppm (in between Cr/Cho and NAA)
Shifted Metabolite Display | 4.7 ppm | We don’t have this option.
# Water Suppressed Shots | 64 | 64
# Non-water-suppressed Shots | 4 | total of 8 reference frames
Save Individual Shots | Yes (raw data and SDAT files) | Yes(raw data saves each individual FID in pfile)
TR (ms) | 5000 | 2000 or 5000 (recommend to avoid T1 weighted)
TE (ms) | 32 or 33 ms | 30ms
Flip angle (deg) | 90 | 90
Shim | Volume based, 2nd Order | auto, first order
Water Suppression | Excitation 100 Hz window | FWHM of the VAPOR pulses is set to 70Hz and is a SLR pulse from CMRR
PNS Mode | High | first level
Gradient Mode | Maximum | first level
TOTAL acquisition time with preparation phases | Roughly 9 minutes | 2:40 for TR=2s or 6:40 for TR=5s **Note: for semiLASER need extra 5 min of setup time

## Slide 9: More to come
I expect that this list of sequences will grow.
Possibilities?
  Myelin water imaging
  Multi-echo fMRI
  Arterial Spin Labelling